## Deviation test 42FIT4011
Flowmeter | Tank ID | Chemical | Injectionpoint | Range 4 mA [kg/t] | Range 20 mA [kg/t] | Density SAS
42FT4011 | 42TB006 | Demulsifier | 2. Step Separator Gjøa | 0 | 5 | 0.91
Accuracy test: | 1 | Note:
Date | 2024-02-24
Flowtransmitter | 42FIT4011
Chemical name | Demulsifier
Last tested | 2019-09-22
1. Firstly warm up the system and stabilize the system with the regulator in manual.
Key settings |  |  | 2. Start the flow comparrison and note every 1 minute the actual kg/hr flow from the Flow Meters.
Setpoint stroke length [%] | 20 |  | 3. Use the Flow Glass as tank if possible.
Regulator [Man/Auto] | Manual |  | 4. Use same filter settings for External and Internal flow meter.
Filter time locally for both [sec] | 10 |  | 5. Reset the totalizer before each test
BEFORE ADJUSTMENTS
Time [minute] | 42FIT4011 [kg/hr] | Extern flowmeter [kg/hr] | Flow glass [ml] | Notes [Faults found] |  |  | Average values for Deviation Test [x] | Test [1] | Test [2] | Test [3] | Date of Test: | 2024-02-24
1 | 0.9261 | 0.5013 | N/A |  |  |  | Setpoint Pump Stroke Length [%] | 20 | 50 | 60 | Flow Transmitter: | 42FIT4011
2 | 0.912 | 0.4988 | N/A | [7 empty] | Flow Measurement Range: | 0 - 5,00 kg/hr
3 | 1.182 | 0.5009 | N/A | Zero calibrated External Flowtransmitter |  |  | 42FIT4011 - FlowMeas [kg/hr] | 0.94794 | 4.7679 | 6.2085 | Last Calibration Date: | 2019-09-22
4 | 1.0416 | 0.501 | N/A |  |  |  | Ext. Flow - FlowActual [kg/hr] | 0.49891 | 2.4267 | 3.1336 | Coriolis Sensor Fabric New | 2016-04-04
5 | 0.9002 | 0.5015 | N/A |  |  |  | Mass Flow Deviation [kg/hr] | 0.44903 | 2.3412 | 3.0748
6 | 0.7725 | 0.49789 | N/A |  |  |  | Mass Flow Deviation [%] | 90.0026 | 96.4779 | 98.1233 | Total Result from all tests
7 | 0.8709 | 0.5012 | N/A | [7 empty] | Average Mass Flow Deviation [%] | 94.8679
8 | 1.1514 | 0.4977 | N/A |  |  |  | 42FIT4011 - TotalizerMeas [kg] | 0.1621 | 0.8022 | 1.059 | Average Totalizer Deviation [%] | 95.4872
9 | 0.8435 | 0.4966 | N/A |  |  |  | Ext. Flow - TotalizerActual [kg] | 0.0859 | 0.4045 | 0.531 | Total Deviation Meas/Actual [%] | 95.1776
10 | 0.8792 | 0.4922 | N/A |  |  |  | Totalizer Deviation [kg] | 0.0762 | 0.3977 | 0.528
[7 empty] | Totalizer Deviation [%] | 88.7078 | 98.3189 | 99.435 | Flow-glass / TotalizerActual Accuracy [%] | 99.9549
Time [minute] | Density Extern flowmeter [kg/m3] | Temp 42FITXXXX [℃] | Temp Extern flowmeter [℃]
1 | 918 | 11.1 | 17.8 |  |  |  | Flow-glass Volume injected [L] | N/A | 0.44 | 0.58 | FlowMeas / FlowActual [%] | 194.8679
2 | [6 empty] | Flow-glass Mass injected [kg] | N/A | 0.40348 | 0.53186 | TotalizerMeas / TotalizerActual [%] | 195.4872
3 | 918 | 11 | 17.8 |  |  |  | Flow-glass/TotalizerActual Deviation [%] | N/A | -0.25216 | 0.16196 | Combined result of FlowMeas and TotalizerMeas [%] | 195.1776
4 | [10 empty] | Total Deviation Ratio Meas/Actual | 1.9518
5 | 918 | 10.9 | 17.9 |  |  |  | 42FIT4011 - Temperature [℃] | 10.92 | 13.66 | 14.44
6 | [6 empty] | Ext. Flow - Temperature [℃] | 17.86 | 16.74 | 16.5
7 | 918 | 10.8 | 17.9 |  |  |  | Ext. Flow - Density [kg/m3] | 918 | 917 | 917
8
9 | 918 | 10.8 | 17.9
10 | [6 empty] | Settings | Before | Calculated
[7 empty] | Flow Calibration Factor [kg/s2] | 284.6641 | 145.8488
Results | 42FITXXXX | Extern flowmeter
Average Mass flow [kg/hr] | 0.94794 | 0.49891 | [5 empty] | Before | After
Average Temperature [℃] | 10.92 | 17.86 |  |  |  |  | Fine adjustments for FCF [kg/s2] | 145.8488 | 141
Average Density [kg/m3] | N/A | 918
[7 empty] | Other Settings | Before | After
[7 empty] | Damping tau Analog Out [sec] | 0 | 10
Note: The totalizer is manually checked on both flowmeters at approximately minute 10. | [6 empty] | Damping filter for Local Display [sec] | 0 | 10
Results | 42FITXXXX | Extern flowmeter |  |  |  |  | Damping Samles [Samples] | 50 | 65
Totalizer [kg] | 0.1621 | 0.0859 |  |  |  |  | Filter Frequency tau [sec] | 0 | 10
Totalizer [L] | 0.17813 | 0.093573 |  |  |  |  | Density Measurement | On | Off
Results Flowglass
Volume delivered from Flowglass [L] | N/A | [5 empty] | AFTER ADJUSTMENTS:
[7 empty] | Average values for Deviation Test [x] | Test [3] | Test [4] |  | Date of Test: | 2024-02-24
[7 empty] | Setpoint Pump Stroke Length [%] | 50 | 20 |  | Flow Transmitter: | 42FIT4011
[11 empty] | Flow Measurement Range: | 0 - 5,00 kg/hr
[7 empty] | 42FIT4011 - Mass Flow [kg/hr] | 2.4438 | 0.5003 |  | Last Calibration Date: | 2019-09-22
[7 empty] | Ext. Flow - Mass Flow [kg/hr] | 2.442 | 0.50097
[7 empty] | Mass Flow Deviation [kg/hr] | 0.00186 | -0.000675 |  | Total Result from all tests
[7 empty] | Mass Flow Deviation [%] | 0.076168 | -0.13474 |  | Average Mass Flow Deviation [%] | -0.029285
[11 empty] | Average Totalizer Deviation [%] | 0.094981
[7 empty] | 42FIT4011 - Totalizer [kg] | 0.41 | 0.084301 |  | Total Deviation Meas/Actual [%] | 0.032848
[7 empty] | Ext. Flow - Totalizer [kg] | 0.409 | 0.084347
[7 empty] | Totalizer Deviation [kg] | 0.001 | -4.6e-05 |  | Flow-glass / TotalizerActual [%] | 98.7838
[7 empty] | Totalizer Deviation [%] | 0.2445 | -0.054537
[11 empty] | FlowMeas / FlowActual [%] | 99.9707
[7 empty] | Flow-glass Volume injected [L] | 0.44 | 0.09 |  | TotalizerMeas / TotalizerActual [%] | 100.095
[7 empty] | Flow-glass Mass injected [kg] | 0.40568 | 0.08298 |  | Combined result of FlowMeas and TotalizerMeas [%] | 100.0328
[7 empty] | 42FIT4011 Flow-glass/Totalizer Deviation [%] | -0.81174 | -1.6207 |  | Total Deviation Ratio Meas/Actual | 1.0003
[7 empty] | 42FIT4011 - Temperature [℃] | 13.14 | 13.14
[7 empty] | Ext. Flow - Temperature [℃] | 13.7 | 13.8
[7 empty] | Ext. Flow - Density [kg/m3] | 922 | 922
Accuracy test: | 2 | Note:
Date | 2024-02-24
Flowtransmitter | 42FIT4011
Chemical name | Demulsifier
Last tested | 2019-09-22
1. Firstly warm up the system and stabilize the system with the regulator in manual.
Key settings |  |  | 2. Start the flow comparrison and note every 1 minute the actual kg/hr flow from the Flow Meters.
Setpoint stroke length [%] | 50 |  | 3. Use the Flow Glass as tank if possible.
Regulator [Man/Auto] | Manual |  | 4. Use same filter settings for External and Internal flow meter.
Filter time locally for both [sec] | 15 |  | 5. Reset the totalizer before each test
Time [minute] | 42FITXXXX [kg/hr] | Extern flowmeter [kg/hr] | Flow glass [ml] | Notes [Faults found] |  |  | Average values for Deviation Test [x]
1 | 4.5981 | 2.4529 | 2200 | Zero Calibrated 42FIT4011 |  |  | Setpoint Pump Stroke Length [%]
2 | 4.6527 | 2.4338
3 | 4.5981 | 2.3818 |  |  |  |  | 42FIT4011 - Mass Flow [kg/hr]
4 | 4.7776 | 2.3229 |  |  |  |  | Ext. Flow - Mass Flow [kg/hr]
5 | 4.8125 | 2.4162 |  |  |  |  | Mass Flow Deviation [kg/hr]
6 | 4.5695 | 2.4504 |  |  |  |  | Mass Flow Deviation [%]
7 | 4.9492 | 2.4532
8 | 4.8653 | 2.46 |  |  |  |  | 42FIT4011 - Totalizer [kg]
9 | 4.9192 | 2.448 |  |  |  |  | Ext. Flow - Totalizer [kg]
10 | 4.9363 | 2.4474 | 2640 |  |  |  | Totalizer Deviation [kg]
Totalizer Deviation [%]
Time [minute] | Density Extern flowmeter [kg/m3] | Temp 42FITXXXX [℃] | Temp Extern flowmeter [℃]
1 | 917 | 13.4 | 16.8 |  |  |  | 42FIT4011 - Temperature [℃]
2 | [6 empty] | Ext. Flow - Temperature [℃]
3 | 917 | 13.6 | 16.8 |  |  |  | Ext. Flow - Density [kg/m3]
4 | [6 empty] | Flow-glass Volume injected [L]
5 | 917 | 13.7 | 16.7 |  |  |  | Flow-glass Mass injected [kg]
6 | [6 empty] | 42FIT4011 Flow-glass/Totalizer Deviation [%]
7 | 917 | 13.8 | 16.7
8
9 | 917 | 13.8 | 16.7
10
Results | 42FITXXXX | Extern flowmeter
Average Mass flow [kg/hr] | 4.7679 | 2.4267
Average Temperature [℃] | 13.66 | 16.74
Average Density [kg/m3] | N/A | 917
Note: The totalizer is manually checked on both flowmeters at approximately minute 10.
Results | 42FITXXXX | Extern flowmeter
Totalizer [kg] | 0.8022 | 0.4045
Totalizer [L] | 0.88154 | 0.44111
Results Flowglass
Volume delivered from Flowglass [L] | 0.44
Accuracy test: | 3
Date | 2024-02-24
Flowtransmitter | 42FIT4011
Chemical name | Demulsifier
Last tested | 2019-09-22
1. Firstly warm up the system and stabilize the system with the regulator in manual.
Key settings |  |  | 2. Start the flow comparrison and note every 1 minute the actual kg/hr flow from the Flow Meters.
Setpoint stroke length [%] | 60 |  | 3. Use the Flow Glass as tank if possible.
Regulator [Man/Auto] | Manual |  | 4. Use same filter settings for External and Internal flow meter.
Filter time locally for both [sec] | 15 |  | 5. Reset the totalizer before each test
Time [minute] | 42FITXXXX [kg/hr] | Extern flowmeter [kg/hr] | Flow glass [ml] | Notes [Faults found] |  |  | Average values for Deviation Test [x]
1 | 6.15 | 3.177 | 2000 |  |  |  | Setpoint Pump Stroke Length [%]
2 | 6.2016 | 3.1202
3 | 6.194 | 3.1313 |  |  |  |  | 42FIT4011 - Mass Flow [kg/hr]
4 | 6.2628 | 3.1344 |  |  |  |  | Ext. Flow - Mass Flow [kg/hr]
5 | 6.2472 | 3.1359 |  |  |  |  | Mass Flow Deviation [kg/hr]
6 | 6.1357 | 3.1141 |  |  |  |  | Mass Flow Deviation [%]
7 | 6.1044 | 3.1313
8 | 6.2728 | 3.1157 |  |  |  |  | 42FIT4011 - Totalizer [kg]
9 | 6.3094 | 3.1388 |  |  |  |  | Ext. Flow - Totalizer [kg]
10 | 6.207 | 3.1378 | 2580 |  |  |  | Totalizer Deviation [kg]
Totalizer Deviation [%]
Time [minute] | Density Extern flowmeter [kg/m3] | Temp 42FITXXXX [℃] | Temp Extern flowmeter [℃]
1 | 917 | 14.2 | 16.5 |  |  |  | 42FIT4011 - Temperature [℃]
2 | [6 empty] | Ext. Flow - Temperature [℃]
3 | 917 | 14.4 | 16.5 |  |  |  | Ext. Flow - Density [kg/m3]
4 | [6 empty] | Flow-glass Volume injected [L]
5 | 917 | 14.5 | 16.5 |  |  |  | Flow-glass Mass injected [kg]
6 | [6 empty] | 42FIT4011 Flow-glass/Totalizer Deviation [%]
7 | 917 | 14.5 | 16.5
8
9 | 917 | 14.6 | 16.5
10
Results | 42FITXXXX | Extern flowmeter
Average Mass flow [kg/hr] | 6.2085 | 3.1336
Average Temperature [℃] | 14.44 | 16.5
Average Density [kg/m3] | N/A | 917
Note: The totalizer is manually checked on both flowmeters at approximately minute 10.
Results | 42FITXXXX | Extern flowmeter
Totalizer [kg] | 1.059 | 0.531
Totalizer [L] | 1.1637 | 0.57906
Results Flowglass
Volume delivered from Flowglass [L] | 0.58
Results of adjusments
Accuracy test: | 4 | Note:
Date | 2024-02-24
Flowtransmitter | 42FIT4011
Chemical name | Demulsifier
Last tested | 2019-09-22
1. Firstly warm up the system and stabilize the system with the regulator in manual.
Key settings |  |  | 2. Start the flow comparrison and note every 1 minute the actual kg/hr flow from the Flow Meters.
Setpoint stroke length [%] | 50 |  | 3. Use the Flow Glass as tank if possible.
Regulator [Man/Auto] | Manual |  | 4. Use same filter settings for External and Internal flow meter.
Filter time locally for both [sec] | 15 |  | 5. Reset the totalizer before each test
Time [minute] | 42FITXXXX [kg/hr] | Extern flowmeter [kg/hr] | Flow glass [ml] | Notes [Faults found] |  |  | Average values for Deviation Test [x]
0 | 2.4177 | 2.4389 | 1670 | Damping for 42FIT: 5 secs --> 60 secs |  |  | Setpoint Pump Stroke Length [%]
0.5 | 2.4215 | 2.4415 |  | Filter Display tau: 10 --> 60 secs
1 | 2.4469 | 2.4411 |  | Damping samples: 50 --> 65 samples |  |  | 42FIT4011 - Mass Flow [kg/hr]
1.5 | 2.447 | 2.443 |  | Filter Frequency tau: 10 --> 20 secs |  |  | Ext. Flow - Mass Flow [kg/hr]
2 | 2.4408 | 2.4367 |  |  |  |  | Mass Flow Deviation [kg/hr]
2.5 | 2.4408 | 2.4367 |  |  |  |  | Mass Flow Deviation [%]
3 | 2.448 | 2.4382 |  | Damping for Ext. Fit: 10 secs --> 20 secs
3.5 | 2.4561 | 2.4433 |  |  |  |  | 42FIT4011 - Totalizer [kg]
4 | 2.4511 | 2.4432 |  |  |  |  | Ext. Flow - Totalizer [kg]
4.5 | 2.4239 | 2.4432 |  |  |  |  | Totalizer Deviation [kg]
5 | 2.4229 | 2.448 |  |  |  |  | Totalizer Deviation [%]
5.5 | 2.4244 | 2.4425
6 | 2.441 | 2.4456 |  |  |  |  | 42FIT4011 - Temperature [℃]
6.5 | 2.46 | 2.4389 |  |  |  |  | Ext. Flow - Temperature [℃]
7 | 2.4612 | 2.4409 |  |  |  |  | Ext. Flow - Density [kg/m3]
7.5 | 2.4477 | 2.4476 |  |  |  |  | Flow-glass Volume injected [L]
8 | 2.4518 | 2.4418 |  |  |  |  | Flow-glass Mass injected [kg]
8.5 | 2.461 | 2.4391 |  |  |  |  | 42FIT4011 Flow-glass/Totalizer Deviation [%]
9 | 2.464 | 2.4428
9.5 | 2.449 | 2.4466 | 2110
Time [minute] | Density Extern flowmeter [kg/m3] | Temp 42FITXXXX [℃] | Temp Extern flowmeter [℃]
1 | 922 | 13 | 13.7
2
3 | 922 | 13 | 13.7
4
5 | 922 | 13.1 | 13.7
6
7 | 922 | 13.3 | 13.7
8
9 | 922 | 13.3 | 13.7
10
Results | 42FITXXXX | Extern flowmeter
Average Mass flow [kg/hr] | 2.4438 | 2.442
Average Temperature [℃] | 13.14 | 13.7
Average Density [kg/m3] | N/A | 922
Note: The totalizer is manually checked on both flowmeters at approximately minute 10.
Results | 42FITXXXX | Extern flowmeter
Totalizer [kg] | 0.41 | 0.409
Totalizer [L] | 0.45055 | 0.4436
Results Flowglass
Volume delivered from Flowglass [L] | 0.44
Accuracy test: | 5 | Note:
Date | 2024-02-24
Flowtransmitter | 42FIT4011
Chemical name | Demulsifier
Last tested | 2019-09-22
1. Firstly warm up the system and stabilize the system with the regulator in manual.
Key settings |  |  | 2. Start the flow comparrison and note every 1 minute the actual kg/hr flow from the Flow Meters.
Setpoint stroke length [%] | 20 |  | 3. Use the Flow Glass as tank if possible.
Regulator [Man/Auto] | Manual |  | 4. Use same filter settings for External and Internal flow meter.
Filter time locally for both [sec] | 15 |  | 5. Reset the totalizer before each test
Time [minute] | 42FITXXXX [kg/hr] | Extern flowmeter [kg/hr] | Flow glass [ml] | Notes [Faults found] |  |  | Average values for Deviation Test [x]
0 | 0.4989 | 0.5048 | 1930 |  |  |  | Setpoint Pump Stroke Length [%]
0.5 | 0.5036 | 0.5012
1 | 0.5005 | 0.5043 |  |  |  |  | 42FIT4011 - Mass Flow [kg/hr]
1.5 | 0.5037 | 0.502 |  |  |  |  | Ext. Flow - Mass Flow [kg/hr]
2 | 0.5018 | 0.503 |  |  |  |  | Mass Flow Deviation [kg/hr]
2.5 | 0.4998 | 0.4976 |  |  |  |  | Mass Flow Deviation [%]
3 | 0.5039 | 0.499
3.5 | 0.4962 | 0.4986 |  |  |  |  | 42FIT4011 - Totalizer [kg]
4 | 0.5018 | 0.4993 |  |  |  |  | Ext. Flow - Totalizer [kg]
4.5 | 0.4986 | 0.4983 |  |  |  |  | Totalizer Deviation [kg]
5 | 0.5016 | 0.5008 |  |  |  |  | Totalizer Deviation [%]
5.5 | 0.4979 | 0.5008
6 | 0.4971 | 0.4978 |  |  |  |  | 42FIT4011 - Temperature [℃]
6.5 | 0.5036 | 0.5032 |  |  |  |  | Ext. Flow - Temperature [℃]
7 | 0.5008 | 0.5023 |  |  |  |  | Ext. Flow - Density [kg/m3]
7.5 | 0.5008 | 0.4992 |  |  |  |  | Flow-glass Volume injected [L]
8 | 0.4993 | 0.501 |  |  |  |  | Flow-glass Mass injected [kg]
8.5 | 0.4972 | 0.4988 |  |  |  |  | 42FIT4011 Flow-glass/Totalizer Deviation [%]
9 | 0.4983 | 0.5028
9.5 | 0.5005 | 0.5046 | 2020
Time [minute] | Density Extern flowmeter [kg/m3] | Temp 42FITXXXX [℃] | Temp Extern flowmeter [℃]
1 | 922 | 13 | 13.8
2
3 | 922 | 13.1 | 13.8
4
5 | 922 | 13.1 | 13.8
6
7 | 922 | 13.2 | 13.8
8
9 | 922 | 13.3 | 13.8
10
Results | 42FITXXXX | Extern flowmeter
Average Mass flow [kg/hr] | 0.5003 | 0.50097
Average Temperature [℃] | 13.14 | 13.8
Average Density [kg/m3] | N/A | 922
Note: The totalizer is manually checked on both flowmeters at approximately minute 10.
Results | 42FITXXXX | Extern flowmeter
Totalizer [kg] | 0.084301 | 0.084347
Totalizer [L] | 0.092638 | 0.091483
Results Flowglass
Volume delivered from Flowglass [L] | 0.09