## Deviation test 42FIT4025
Flowmeter | Tank ID | Chemical | Injectionpoint | Range 4 mA [kg/hr] | Range 20 mA [kg/hr] | Density SAS
42FT4025 | 42TB054 | pH Stabilisator | Vega MEG lagertank | 0 | 58.7 | 1.33
Accuracy test: | 1
Date | 2024-02-23
Flowtransmitter | 42FIT4025
Chemical name | pH Stabilizer
Last tested | 2023-10-23
1. Firstly warm up the system and stabilize the system with the regulator in manual.
Key settings |  |  | 2. Start the flow comparrison and note every 1 minute the actual kg/hr flow from the Flow Meters.
Setpoint stroke length [%] | 44 |  | 3. Use the Flow Glass as tank if possible.
Regulator [Man/Auto] | Manual |  | 4. Use same filter settings for External and Internal flow meter.
Filter time locally for both [sec] | 15 |  | 5. Reset the totalizer before each test
BEFORE ADJUSTMENTS
Time [minutes] | 42FITXXXX [kg/hr] | Extern flowmeter [kg/hr] | Flow glass [ml] | Notes [Faults found] |  |  | Average values for Deviation Test [x] | Test [1] | Test [2] | Date of Test: | 2024-02-23
0-1 minute | 19.6 | 19.711 | N/A | Pump 42PD054B is leaking and using oil |  |  | Setpoint Pump Stroke Length [%] | 44 | 52 | Flow Transmitter: | 42FIT4025
1-2 minute | 19.82 | 19.843 | N/A | 42BL4629 Leaking through stem (Event made) | [5 empty] | Flow Measurement Range: | 0 - 58,70 kg/hr
2-3 minute | 20.087 | 20.018 | N/A | 42BL4598 Leaking (Replaced 23.02.24) |  |  | 42FIT4025 - FlowMeas [kg/hr] | 19.9919 | 17.6 | Last Calibration Date: | 2023-10-23
3-4 minute | 20.114 | 20.136 | N/A |  |  |  | Ext. Flow - FlowActual [kg/hr] | 20.2331 | 17.64 | Coriolis Sensor Fabric New | 2023-09-01
4-5 minute | 20.13 | 20.22 | N/A |  |  |  | Mass Flow Deviation [kg/hr] | -0.2412 | -0.04
5-6 minute | 19.78 | 20.348 | N/A | Zero calibrated External Flowtransmitter |  |  | Mass Flow Deviation [%] | -1.1921 | -0.22676 | Total Result from all tests
6-7 minute | 20.246 | 20.38 | N/A | [6 empty] | Average Mass Flow Deviation [%] | -0.70943
7-8 minute | 19.2 | 20.47 | N/A |  |  |  | 42FIT4025 - TotalizerMeas [kg] | 3.3958 | 4.8234 | Average Totalizer Deviation [%] | -0.63016
8-9 minute | 20.47 | 20.574 | N/A |  |  |  | Ext. Flow - TotalizerActual [kg] | 3.4205 | 4.8495 | Total Deviation Meas/Actual [%] | -0.66979
9-10 minute | 20.472 | 20.631 | N/A |  |  |  | Totalizer Deviation [kg] | -0.0247 | -0.0261
[7 empty] | Totalizer Deviation [%] | -0.72212 | -0.5382 | FlowMeas / FlowActual [%] | 99.2906
Time [minutes] | Density Extern flowmeter [kg/m3] | Temp 42FITXXXX [℃] | Temp Extern flowmeter [℃] | [6 empty] | TotalizerMeas / TotalizerActual [%] | 99.3698
0-1 minute | 1320 | 17.4 | 18.6 |  |  |  | 42FIT4025 - Temperature [℃] | 17.4 | 17.6 | Combined result of FlowMeas and TotalizerMeas [%] | 99.3302
1-2 minute | [6 empty] | Ext. Flow - Temperature [℃] | 18.22 | 17.64 | Total Deviation Ratio Meas/Actual | 0.9933
2-3 minute | 1321 | 17.3 | 18.4 |  |  |  | Ext. Flow - Density [kg/m3] | 1321.6 | 1322.8
3-4 minute
4-5 minute | 1321 | 17.4 | 18.1
5-6 minute
6-7 minute | [6 empty] | No Adjustments done
7-8 minute | 1322 | 17.5 | 18
8-9 minute
9-10 minute | 1324 | 17.4 | 18
Results | 42FITXXXX | Extern flowmeter
Average Mass flow [kg/hr] | 19.9919 | 20.2331
Average Temperature [℃] | 17.4 | 18.22
Average Density [kg/m3] | N/A | 1321.6
Note: The totalizer is manually checked on both flowmeters at approximately minute 10.
Results | 42FITXXXX | Extern flowmeter
Totalizer [kg] | 3.3958 | 3.4205
Totalizer [L] | 2.5532 | 2.5882
Results Flowglass
Volume delivered from Flowglass [L] | N/A
Accuracy test: | 2 | Note:
Date | 2024-02-23
Flowtransmitter | 42FIT4025
Chemical name | pH Stabilizer
Last tested | 2023-10-23
1. Firstly warm up the system and stabilize the system with the regulator in manual.
Key settings |  |  | 2. Start the flow comparrison and note every 1 minute the actual kg/hr flow from the Flow Meters.
Setpoint stroke length [%] | 52 |  | 3. Use the Flow Glass as tank if possible.
Regulator [Man/Auto] | Manual |  | 4. Use same filter settings for External and Internal flow meter.
Filter time locally for both [sec] | 15 |  | 5. Reset the totalizer before each test
Time [minutes] | 42FITXXXX [kg/hr] | Extern flowmeter [kg/hr] | Flow glass [ml] | Notes [Faults found]
0-1 minute | 28.064 | 28.28 | N/A | Due to pump leakage and a accurate measurement, only 1 deviation test was performed
1-2 minute | 27.884 | 28.37 | N/A | Zero-Calibration with operating conditions and settings verification was done before this test
2-3 minute | 28.408 | 28.46 | N/A
3-4 minute | 27.09 | 28.54 | N/A | Zero calibrated 42FIT4025
4-5 minute | 28.48 | 28.58 | N/A
5-6 minute | 28.6 | 28.68 | N/A
6-7 minute | 28.725 | 28.69 | N/A
7-8 minute | 28.611 | 28.75 | N/A
8-9 minute | 28.67 | 28.76 | N/A
9-10 minute | 28.59 | 28.81 | N/A | Conclusion: Already installed new coriolis and calibrated in 10.2023, was a good test to test my external sensor.
Time [minutes] | Density Extern flowmeter [kg/m3] | Temp 42FITXXXX [℃] | Temp Extern flowmeter [℃]
0-1 minute | 1324 | 17.8 | 17.8
1-2 minute
2-3 minute | 1323 | 17.7 | 17.8
3-4 minute
4-5 minute | 1321 | 17.5 | 17.5
5-6 minute
6-7 minute
7-8 minute | 1322 | 17.5 | 17.6
8-9 minute
9-10 minute | 1324 | 17.5 | 17.5
Results | 42FITXXXX | Extern flowmeter
Average Mass flow [kg/hr] | 28.3122 | 28.592
Average Temperature [℃] | 17.6 | 17.64
Average Density [kg/m3] | N/A | 1322.8
Note: The totalizer is manually checked on both flowmeters at approximately minute 10.
Results | 42FITXXXX | Extern flowmeter
Totalizer [kg] | 4.8234 | 4.8495
Totalizer [L] | 3.6266 | 3.6661
Results Flowglass
Volume delivered from Flowglass [L] | N/A